## Slide 1: Development and Performance Analysis of a Novel Pulse Oximetry Measurement Circuit
Master’s Thesis Mikko Tuohimaa

## Slide 2: Contents
Physiology of Oxygenation
Introduction to Pulse Oximetry
The New Measurement Circuit
Performance Analysis
Conclusions

## Slide 3: Tissue Oxygenation
Proper oxygenation vital for all organs
Oxygen doesn’t dissolve easily in blood plasma
Specialized protein complex, hemoglobin, in red blood cells able to reversibly bind oxygen molecules

## Slide 4: Hemoglobin
Hb: 97% of the dry content of red blood cells

## Slide 5: Photoplethysmography
Measuring volume with light. Basically, measures light absorption due to subject
thickness and
composition. Each substance i has its own absorption coefficient ελ for each wavelength λ:
  -> Use several wavelengths to find out concentrations ci

## Slide 6: Hemoglobin Absorption Spectra

## Slide 7: Photoplethysmography
Optical path length l changes due to heart beating: A time-variant signal
Problem: extract information about arterial blood
Solution: focus on the AC signal

## Slide 8: Pulse Oximetry: Theory
Modulation ratio:
Ratio of Ratios:
Oxygen Saturation:

## Slide 9: Pulse Oximetry: Practice
Red and IR LEDs on one side of a tissue bed, a photodetector on the other side
A pulsed sampling scheme with ambient removal

## Slide 10: Pulse Oximetry: Sampling Timing

## Slide 11: Pulse Oximetry: Sources of Error
System Noise
Motion Artifacts
Medical Dyes in Blood
High dysfunctional hemoglobin content (e.g. CO poisoning)

## Slide 12: The New Measurement Circuit
Designed to be an autonomous module for easy integration

## Slide 13: The New Measurement Circuit
The measurement hardware controlled by an embedded microcontroller
Control parameters:
Timing
Receiver gain and filtering properties
Transmitter LED current
Diagnostics

## Slide 14: The New Measurement Circuit: Tx
A switched constant current source with a reference provided by an 8-bit DAC

## Slide 15: The New Measurement Circuit: Rx
A differential transimpedance amplifier with 7 selectable gains per channel feeding 4 sample-and-hold circuits and a 22-bit ΣΔ ADC.

## Slide 16: Performance Analysis: Methods
Key performance indicator: measurement noise
Compact design and extremely small signals mean that it’s impossible to measure subsystems directly
The performance of individual subsystems can be examined with regression analysis
One isolated system parameter is modified at a time and the resulting change in performance is fitted into the system model

## Slide 17: Performance Analysis: Data
Isolated receiver operating parameters:
Gain
Sampling time and low-pass cutoff freq
Conversion time
Signal level
The transmitter was found satisfactory
-> main focus on the receiver.

## Slide 18: Performance Analysis: Data
Design Spec:
σ < 18
Performance OK with zero signal

## Slide 19: Performance Analysis: Data
Pulsed current introduced
-> performance degrades
No unwanted crosstalk

## Slide 20: Performance Analysis: Data

## Slide 21: Performance Analysis: Data

## Slide 22: Performance Analysis: Results 1
Optimal ambient scaling is ca. 0.7, yet signal paths for red and red ambient have no difference in gain.
Conclusion: there’s a significant amount of seemingly uncorrelated noise in the two signals. Transmitter has been ruled out as the source.

## Slide 23: Performance Analysis: Results 2
There is a component in the signal chain the noise behavior of which is signal-dependent. The buffer feeding the ADC was identified as the source.

## Slide 24: Performance Analysis: Results 3
A temporary fix, while not perfect, proved the hypothesis true.

## Slide 25: Conclusions
Bug identified in the receiver signal chain
No expensive precision measurement equipment needed
Performance deemed satisfactory provided that the bug is adequately fixed